Zuzlewski Stéphane August 2001

hardware

tracking

schema

Volume III Version 1.5.1

Table of contents

1.- Introduction. ..........................................................…………………………...	2

2.- Database structure. ........................................…………………………..........	3

3.- Database relations. ...................................…………………………................	12

4.- Database attributes. .......................…………………………..........................	22

1.- Introduction :

This paper is about the design of the NCEDC schema for the hardware tracking. This schema was designed to be able to track any station from the sensor to the datalogger by their serial numbers. It can also generate automatically the instrument responses schema.

There are 28 relations in the schema. The information stored in these relations includes:

- Responses description (high-pass, low-pass, FIR filters; Poles & zeros; Polynomial).

- Sensors and components information.

- Analog filters, amplifiers information.

- Dataloggers, digitizers information.

- Dataloggers logical channels information.

- Tracking information.

The remainder of this volume consists of four chapters:

Chapter 2 		Database structure

Each relation is defined.

Chapter 3		Database relations

Each relation is described to identify the key fields and the links among the relations.

Chapter 4		Database attributes

Each attribute is described.

2.- Database structure :

This chapter defines the physical structure of each table, as it exists within the DBMS data dictionary. The name of the relation appears in bold print at the top left and a description of the table at the top right. Bold attributes represent primary keys, italic attributes represent foreign keys, and underlined attributes stand for CSS attributes.

| Response Summary information on a sequence of responses |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| seqresp_id | int | Sequence of responses identifier |
| resp_nb | int | Response number |
| resp_type | char(1) | Response type |
| resp_id | int | Response identifier |
| unit_in | int | Signal input units |
| unit_out | int | Signal output units |
| r_type | char(1) | SEED response type |
| lddate | date | Load date |

| Response_HP Summary information on a high-pass filter |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| hp_id | int | Filter identifier |
| filter_type | char(2) | Filter type |
| nb_pole | int | Number of poles |
| corner_freq | float | Corner frequency of filter |
| damping_value | float | Fraction of critical damping |
| lddate | date | Load date |

| Response_LP Summary information on a low-pass filter |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| lp_id | int | Filter identifier |
| filter_type | char(2) | Filter type |
| nb_pole | int | Number of poles |
| corner_freq | float | Corner frequency of filter |
| damping_value | float | Fraction of critical damping |
| lddate | date | Load date |

| Response_PZ Summary information on poles & zeros |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| pz_id | int | Poles & zeros identifier |
| pz_nb | int | Pole/zero number |
| type | char(1) | Pole/Zero |
| r_value | float | Real pole/zero |
| r_error | float | Real pole/zero error |
| i_value | float | Imaginary pole/zero |
| i_error | float | Imaginary pole/zero error |
| lddate | date | Load date |

| Response_PN Summary information on a polynomial |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| pn_id | int | Polynomial identifier |
| name | char(80) | Polynomial name |
| poly_type | char(1) | Type of polynomial approximation |
| lower_bound | float | Polynomial lower bound |
| upper_bound | float | Polynomial upper bound |
| max_error | float | Maximum error of polynomial approx. |
| nb_coeff | int | Number of polynomial coefficients |
| lddate | date | Load date |

| Response_PN_Data Coefficients of a polynomial |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| pn_id | int | Polynomial identifier |
| pn_nb | int | Coefficient number |
| pn_value | float | Coefficient value |

| Sensor Summary information on a sensor |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sensor_id | int | Sensor identifier |
| name | char(80) | Sensor name |
| serial_nb | char(80) | Sensor serial number |
| ondate | date | Start date |
| offdate | date | Off date |
| nb_component | int | Number of components |
| lddate | date | Load date |

| Sensor_Component Component of a sensor |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sensor_id | int | Sensor identifier |
| component_nb | int | Component number |
| channel_comp | char(2) | Channel |
| component_type | char(1) | Type |
| sensitivity | float | Sensitivity |
| frequency | float | Frequency |
| seqresp_id | int | Sequence of responses identifier |
| lddate | date | Load date |

| Filamp Summary information on a filter-amplifier |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| filamp_id | int | Filter-amplifier identifier |
| name | char(80) | Filter-amplifier name |
| serial_nb | char(80) | Filter-amplifier serial number |
| ondate | date | Start date |
| offdate | date | Off date |
| nb_pchannel | int | Number of physical channels |
| lddate | date | Load date |

| Filamp_PChannel Physical channel of a filter-amplifier |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| filamp_id | int | Filter-amplifier identifier |
| pchannel_nb | int | Physical channel number |
| gain | float | Gain |
| frequency | float | Calibration frequency |
| seqresp_id | int | Sequence of responses identifier |
| lddate | date | Load date |

| Datalogger Summary information on a datalogger |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| data_id | int | Datalogger identifier |
| data_type | char(80) | Datalogger type |
| serial_nb | char(80) | Datalogger serial number |
| firmware_nb | char(80) | Datalogger firmware number |
| software | char(80) | Software (ushear, mshear, …) |
| software_nb | char(80) | Software version number |
| ondate | date | Start date |
| offdate | date | Off date |
| nb_board | int | Number of boards |
| word_32 | int | 32 bit word order |
| word_16 | int | 16 bit word order |
| lddate | date | Load date |

| Datalogger_Board Board of a datalogger |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| data_id | int | Datalogger identifier |
| board_nb | int | Board number |
| serial_nb | char(80) | Board serial number |
| nb_module | int | Number of modules |
| lddate | date | Load date |

| Datalogger_Module Module of a datalogger’s board |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| data_id | int | Datalogger identifier |
| board_nb | int | Board number |
| module_nb | int | Module number |
| serial_nb | char(80) | Module serial number |
| sensitivity | float | Sensitivity |
| lddate | date | Load date |

| Filter_FIR Summary information on a FIR filter |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| fir_id | int | FIR filter identifier |
| name | char(80) | Filter name |
| symmetry | char(1) | Filter symmetry |
| gain | float | Filter gain |
| lddate | date | Load date |

| Filter_FIR_Data Coefficients of a FIR filter |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| fir_id | int | FIR filter identifier |
| coeff_nb | int | Numerator/denominator number |
| type | char(1) | Numerator/Denominator |
| coefficient | float | Numerator/denominator coefficient |
| error | float | Numerator/denominator error |

| Filter Summary information on a filter |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| filter_id | int | Filter identifier |
| gain | float | Filter gain |
| frequency | float | Frequency for filter gain |
| in_sp_rate | float | Input sample rate |
| out_sp_rate | float | Output sample rate |
| offset | int | Decimation offset |
| delay | float | Estimated delay |
| correction | float | Correction applied |
| seqresp_id | int | Sequence of responses identifier |
| lddate | date | Load date |

| Filter_Sequence Summary information on a sequence of filters |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| seqfil_id | int | Sequence of filters identifier |
| name | char(32) | Sequence of filters name |
| nb_filter | int | Number of filters |
| gain | float | Overall gain |
| frequency | float | Frequency |
| lddate | date | Load date |

| Filter_Sequence_Data Filter of a sequence |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| seqfil_id | int | Sequence of filters identifier |
| filter_nb | int | Filter number |
| filter_id | int | Filter identifier |

| Station Summary information on a station |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| lat | float | Latitude |
| lon | float | Longitude |
| elev | float | Elevation |
| staname | char(50) | Station description |
| nb_sensor | int | Number of sensors |
| nb_filamp | int | Number of filter-amplifiers |
| nb_digi | int | Number of digitizers |
| nb_data | int | Number of dataloggers |
| datumhor | char(8) | Datum type for latitude and longitude |
| datumver | char(8) | Datum type for vertical |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Sensor Data associating a station with a sensor |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| sensor_nb | int | Sensor number |
| sensor_id | int | Sensor identifier |
| lat | float | Latitude |
| lon | float | Longitude |
| elev | float | Elevation |
| edepth | float | Emplacement depth |
| nb_component | int | Number of components |
| datumhor | char(8) | Datum type for latitude and longitude |
| datumver | char(8) | Datum type for vertical |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Sensor_Component Data associating physical channels |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| sensor_nb | int | Sensor number |
| component_nb | int | Component number |
| next_hard_type | char(1) | Next hardware type |
| next_hard_nb | int | Next hardware number |
| next_hard_pchannel | int | Next hardware physical channel |
| azimuth | float | Azimuth |
| dip | float | Dip |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Filamp Data associating a station with a filter-amplifier |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| filamp_nb | int | Filter-amplifier number |
| filamp_id | int | Filter-amplifier identifier |
| nb_pchannel | int | Number of physical channels |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Filamp_PChannel Data associating physical channels |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| filamp_nb | int | Filter-amplifier number |
| pchannel_nb | int | Physical channel number |
| next_hard_type | char(1) | Next hardware type |
| next_hard_nb | int | Next hardware number |
| next_hard_pchannel | int | Next hardware physical channel |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Digitizer Data associating a station with a digitizer |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| digi_nb | int | Digitizer number |
| serial_nb | char(80) | Digitizer serial number |
| nb_pri_pchannel | int | Number of primary physical channels |
| nb_aux_pchannel | int | Number of auxiliary physical channels |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Digitizer_PChannel Data associating physical channels |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| digi_nb | int | Digitizer number |
| pchannel_nb | int | Physical channel number |
| data_nb | int | Datalogger number |
| data_pchannel | int | Datalogger physical channel |
| digi_type | char(3) | Digitizer type |
| digi_polarity | char(1) | Digitizer polarity |
| digi_channel | int | Digitizer channel |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Datalogger Data associating a station with a datalogger |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| data_nb | int | Datalogger number |
| data_id | int | Datalogger identifier |
| nb_pchannel | int | Number of physical channels |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Datalogger_PChannel Data associating channels |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| data_nb | int | Datalogger number |
| pchannel_nb | int | Physical channel number |
| board_type | char(1) | Board type |
| channel_type | char(1) | Channel type |
| seed_io | char(2) | SEED instrument and orientation |
| nb_lchannel | int | Number of logical channels |
| ondate | date | Start date |
| offdate | date | Off date |
| lddate | date | Load date |

| Station_Datalogger_LChannel Summary information on a channel |  |  |
| --- | --- | --- |
| Attribute | Type | Description |
| sta | char(6) | Station code |
| net | char(8) | Network code |
| data_nb | int | Datalogger number |
| pchannel_nb | int | Physical channel number |
| lchannel_nb | char(2) | Logical channel number |
| seqfil_id | int | Sequence of filters identifier |
| seedchan | char(3) | SEED channel name |
| channel | char(3) | Channel name |
| channelsrc | char(8) | Domain of channel name |
| location | char(2) | SEED location |
| rgain | float | Gain at the response frequency |
| rfrequency | float | Frequency for response |
| samprate | float | Sampling rate |
| clock_drift | float | Maximum clock drift |
| flags | char(27) | SEED channel flags |
| data_format | char(80) | Data format |
| comp_type | int | Compression type |
| unit_signal | int | Units of signal responses |
| unit_calib | int | Units of calibration |
| block_size | int | Block size |
| ondate | date | Start date |
| offdate | date | Off date |
| remark | char(30) | Optional comment |
| lddate | date | Load date |

3.- Database relations :

This chapter describes the DBMS relations that comprise the version 1.5.1 schema. The information given here, along with that in Chapter 4, Database attributes, constitutes the data dictionary. There’s an entry for each relation. Within the entry, the relation’s name appears first, followed by a list of its attributes. A brief description completes the entry.

Name: Response Keys: Primary: seqresp_id, resp_nb Foreign: resp_id, unit_in, unit_out Data: resp_type, r_type, lddate Description: This relation provides information about a sequence of responses.

Name: Response_HP Keys: Primary: hp_id Foreign: Data: filter_type, nb_pole, corner_freq, damping_value, lddate Description: This relation describes a response defined by a high-pass filter.

Name: Response_LP Keys: Primary: lp_id Foreign: Data: filter_type, nb_pole, corner_freq, damping_value, lddate Description: This relation describes a response defined by a low-pass filter.

Name: Response_PZ Keys: Primary: pz_id, pz_nb, type Foreign: Data: r_value, r_error, i_value, i_error, lddate Description: This relation describes a response defined by poles and zeros.

Name: Response_PN Keys: Primary: pn_id Foreign: Data: name, poly_type, lower_bound, upper_bound, max_error, nb_coeff, lddate Description: This relation describes a response defined by a polynomial.

Name: Response_PN_Data Keys: Primary: pn_id, pn_nb Foreign: Data: pn_value Description: This relation provides information about the coefficients of a polynomial.

Name: Sensor Keys: Primary: sensor_id Foreign: Data: name, serial_nb, ondate , offdate, nb_component, lddate Description: This relation describes a sensor.

Name: Sensor_Component Keys: Primary: sensor_id, component_nb Foreign: seqresp_id Data: channel_comp, component_type, sensitivity, frequency, lddate Description: This relation describes a component of a sensor.

Name: Filamp Keys: Primary: filamp_id Foreign: Data: name, serial_nb, ondate , offdate, nb_pchannel, lddate Description: This relation describes a filter-amplifier.

Name: Filamp_PChannel Keys: Primary: filamp_id, pchannel_nb, frequency Foreign: seqresp_id Data: gain, lddate Description: This relation describes a physical channel of a filter-amplifier.

Name: Datalogger Keys: Primary: data_id Foreign: Data: data_type, serial_nb, firmware_nb, software, software_nb, ondate, offdate, nb_board, word_32, word_16, lddate Description: This relation describes a datalogger.

Name: Datalogger_Board Keys: Primary: data_id, board_nb Foreign: Data: serial_nb, nb_module, lddate Description: This relation describes a board of a datalogger.

Name: Datalogger_Module Keys: Primary: data_id, board_nb, module_nb Foreign: Data: serial_nb, sensitivity, lddate Description: This relation describes a module of a datalogger’s board.

Name: Filter_FIR Keys: Primary: fir_id Foreign: Data: name, symmetry, gain, lddate Description: This relation describes a FIR filter.

Name: Filter_FIR_Data Keys: Primary: fir_id, coeff_nb Foreign: Data: type, coefficient, error Description: This relation provides information about the coefficients of a FIR filter.

Name: Filter Keys: Primary: filter_id Foreign: seqresp_id Data: gain, frequency, in_sp_rate, out_sp_rate, offset, delay, correction, lddate Description: This relation describes a filter.

Name: Filter_Sequence Keys: Primary: seqfil_id Foreign: Data: name, nb_filter, gain, frequency, lddate Description: This relation describes a sequence of filters.

Name: Filter_Sequence_Data Keys: Primary: seqfil_id, filter_nb Foreign: filter_id Data: Description: This relation provides information about the filters of a sequence.

Name: Station Keys: Primary: sta, net, ondate Foreign: Data: lat, lon, elev, staname, nb_sensor, nb_filamp, nb_digi, nb_data, datumhor, datumver, offdate, lddate Description: This relation describes and names a point of the earth where seismic measurements are made (e.g. the location of a seismic instrument or array).

Name: Station_Sensor Keys: Primary: sta, net, sensor_nb, ondate Foreign: sensor_id Data: lat, lon, elev, edepth, nb_component, datumhor, datumver, offdate, lddate Description: This relation describes a sensor for a particular station.

Name: Station_Sensor_Component Keys: Primary: sta, net, sensor_nb, component_nb, ondate Foreign: Data: next_hard_type, next_hard_nb, next_hard_pchannel, azimuth, dip, offdate, lddate Description: This relation describes a component of a sensor for a particular station.

Name: Station_Filamp Keys: Primary: sta, net, filamp_nb, ondate Foreign: filamp_id Data: nb_pchannel, offdate, lddate Description: This relation describes a filter-amplifier for a particular station.

Name: Station_Filamp_PChannel Keys: Primary: sta, net, filamp_nb, pchannel_nb, ondate Foreign: Data: next_hard_type, next_hard_nb, next_hard_pchannel, offdate, lddate Description: This relation describes a physical channel of a filter-amplifier for a particular station.

Name: Station_Digitizer Keys: Primary: sta, net, digi_nb, ondate Foreign: Data: serial_nb, nb_pri_channel, nb_aux_channel, offdate, lddate Description: This relation describes a digitizer for a particular station.

Name: Station_Digitizer_PChannel Keys: Primary: sta, net, digi_nb, pchannel_nb, ondate Foreign: Data: data_nb, data_pchannel, digi_type, digi_polarity, digi_channel, offdate, lddate Description: This relation describes a physical channel of a digitizer for a particular station.

Name: Station_Datalogger Keys: Primary: sta, net, data_nb, ondate Foreign: data_id Data: nb_pchannel, offdate, lddate Description: This relation describes a datalogger for a particular station.

Name: Station_Datalogger_PChannel Keys: Primary: sta, net, data_nb, pchannel_nb, ondate Foreign: Data: board_type, channel_type, seed_io, nb_lchannel, offdate, lddate Description: This relation describes a physical channel of a datalogger for a particular station.

Name: Station_Datalogger_LChannel Keys: Primary: sta, net, data_nb, pchannel_nb, lchannel_nb, ondate Foreign: seqfil_id, comp_type, unit_signal, unit_calib Data: seedchan, channel, channelsrc, location, rgain, rfrequency, samprate, clock_drift, flags, data_format, block_size, offdate, remark, lddate Description: This relation describes and names a particular channel for a particular station.

4.- Database attributes :

This chapter describes each of the attributes used in the version 1.5.1 schema. Descriptions of the relations are found in Chapter 3, Database relations.

Name: azimuth Relations: Station_Sensor_Component Description: The azimuth of the instrument from north, clockwise. Datatype: float NULL Value: Yes Range: 0.0  azimuth  360.0 Units: Degrees

Name: block_size Relations: Station_Datalogger_LChannel Description: The record length of the data. Datatype: int NULL Value: No Range: 256  block_size  4096

Name: board_nb Relations: Datalogger_Board, Datalogger_Module Description: This attribute represents the board number of a datalogger. Datatype: int NULL Value: No Range: board_nb  1

Name: board_type Relations: Station_Datalogger_PChannel Description: This attribute describes the board type on the datalogger where the physical channel resides. Regular seismic channels are usually on a primary board. Auxiliary channels (internal temperature, external power, …) are usually on an auxiliary board. A value of external is used to represent a digitizer that is not part of the datalogger (for example an embedded digitizer in a digital pressure sensor). A digital board is used to represent truly digital sensors which don’t actually have any digitizing stage from Volts to Counts (for example a rainfall sensor that uses a digital counter). Datatype: char(1) NULL Value: No Range: board_type  { P, A, E, D } P = primary A = auxiliary E = external D = digital

Name: channel Relations: Station_Datalogger_LChannel Description: Channel name. Datatype: char(3) NULL Value: Yes Range: Any valid channel name, which may be network-specific. Networks using SEED channel names should use the same value for channel and seedchan

Name: channel_comp Relations: Sensor_Component Description: Component channel. Datatype: char(2) NULL Value: Yes

Name: channel_type Relations: Station_Datalogger_PChannel Description: This attribute describes the type of the input channel. An input channel is usually classified as physical when the data is digitized (velocity, acceleration, …). An input channel is defined as synthetic when the data is not derived directly from a sensor (clock frequency, available telemetry buffer channel, clock quality, …). Datatype: char(1) NULL Value: No Range: channel_type  { P, S } P = physical S = synthetic

Name: channelsrc Relations: Station_Datalogger_LChannel Description: Domain for channel. This specifies what naming convention is used for the channel name (i.e. SEED, USGS, etc.). Datatype: char(8) NULL Value: Yes Range: Any free-format string up to 8 characters long

Name: clock_drift Relations: Station_Datalogger_LChannel Description: A tolerance value, measured in seconds per sample, used as a thresold for time error detection in the data. The number of samples in a record are multiplied by this value to calculate a maximum drift allowed for the time in the next record. If the difference in time is less than this drift, consider them in the same time series. Datatype: float NULL Value: Yes Range: clock_drift  0.0 Units: Seconds

Name: coefficient Relations: Filter_FIR_Data Description: The numerator/denominator coefficient value. Datatype: float NULL Value: No

Name: coeff_nb Relations: Filter_FIR_Data Description: The numerator/denominator number. Datatype: int NULL Value: No Range: coeff_nb  1

Name: comp_type Relations: Station_Datalogger_LChannel Description: Compression type. A lookup key that refers to the Format Dictionary. Datatype: int NULL Value: No

Name: component_nb Relations: Sensor_Component, Station_Sensor_Component Description: This attribute represents the component number. Datatype: int NULL Value: No Range: component_nb  1

Name: component_type Relations: Sensor_Component Description: This attribute denotes the component type. Datatype: char(1) NULL Value: Yes

Name: corner_freq Relations: Response_HP, Response_LP Description: This attribute represents the corner frequency of a high/low-pass filter. Datatype: float NULL Value: No Range: corner_freq > 0.0 Units: Hertz

Name: correction Relations: Filter Description: The time shift applied to the time tag due to delay at this stage of the filter; a negative number indicating the amount of time added to the former time tag. The actual delay is difficult to estimate, and the correction applied neglects dispersion. This field allows to know how much correction was used, in case a more accurate correction is to be applied later. A zero here implies no correction was done. Datatype: float NULL Value: No

Name: damping_value Relations: Response_HP, Response_LP Description: This attribute represents the fraction of critical damping of a high/low-pass filter. Datatype: float NULL Value: No

Name: data_format Relations: Station_Datalogger_LChannel Description: This attribute represents the data format. Datatype: char(80) NULL Value: No Range: data_format  { SEED ¦ …} SEED = SEED format

Name: data_id Relations: Datalogger, Datalogger_Board, Datalogger_Module, Station_Datalogger Description: Datalogger identifier. Each datalogger is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: data_nb Relations: Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: This attribute represents the datalogger number. Datatype: int NULL Value: No Range: data_nb  1

Name: data_pchannel Relations: Station_Digitizer_PChannel Description: This attribute represents a datalogger physical channel. Datatype: int NULL Value: No Range: data_pchannel  1

Name: data_type Relations: Datalogger Description: This attribute represents a datalogger type. Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters long

Name: datumhor Relations: Station, Station_Sensor Description: Datum type for horizontal (latitude and longitude) coordinates. Datatype: char(8) NULL Value: Yes Range: datumhor  { NAD27 ¦ WGS84}

Name: datumver Relations: Station, Station_Sensor Description: Datum type for depth or elevation. Datatype: char(8) NULL Value: Yes Range: datumver  { NAD27 ¦ WGS84 ¦ AVERAGE } AVERAGE = average of stations

Name: delay Relations: Filter Description: The estimated pure delay for the filter. Datatype: float NULL Value: Yes Units: Seconds

Name: digi_channel Relations: Station_Digitizer_PChannel Description: This attribute represents a digitizer module on a board. Datatype: int NULL Value: No Range: digi_channel  1

Name: digi_nb Relations: Station_Digitizer, Station_Digitizer_PChannel Description: This attribute represents the digitizer number. Datatype: int NULL Value: No Range: digi_nb  1

Name: digi_polarity Relations: Station_Digitizer_PChannel Description: This attribute represents the digitizer polarity. Datatype: char(1) Null Value: No Range: digi_polarity  { N ¦ … }

Name: digi_type Relations: Station_Digitizer_PChannel Description: This attribute represents the digitizer type. Datatype: char(3) NULL Value: No Range: digi_type  { DSP ¦ AUX }

Name: dip Relations: Station_Sensor_Component Description: The dip of the instrument, down from horizontal. Datatype: float NULL Value: Yes Range: -90.0  dip  90.0 Units: Degrees

Name: edepth Relations: Station_Sensor Description: Emplacement depth. This attribute gives the depth at which the instrument is positioned, relative to the value of elev. Datatype: float NULL Value: Yes Range: edepth  0.0 Units: Kilometers

Name: elev Relations: Station, Station_Sensor Description: Elevation. This attribute is the elevation of a seismic station relative to mean sea level. Datatype: float NULL Value: Yes Range: -10.0  elev  10.0 Units: Kilometers

Name: error Relations: Filter_FIR_Data Description: This attributes denotes the numerator/denominator error. Datatype: float NULL Value: Yes

Name: filamp_id Relations: Filamp, Filamp_PChannel, Station_Filamp Description: Filter-amplifier identifier. Each filter-amplifier is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: filamp_nb Relations: Station_Filamp, Station_Filamp_PChannel Description: This attribute represents the filter-amplifier number. Datatype: int NULL Value: No Range: filamp_nb  1

Name: filter_id Relations: Filter, Filter_Sequence_Data Description: Filter identifier. Each filter is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: filter_nb Relations: Filter_Sequence_Data Description: This attribute represents the filter number in a sequence. Datatype: int NULL Value: No Range: filter_nb  1

Name: filter_type Relations: Response_HP, Response_LP Description: This attribute denotes the type of a high/low-pass filter. Datatype: char(2) NULL Value: No Range: filter_type  { BW ¦ DG ¦ ND } BW = butterworth DG = damping given ND = no damping

Name: fir_id Relations: Filter_FIR, Filter_FIR_Data Description: FIR Filter identifier. Each FIR filter is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: firmware_nb Relations: Datalogger Description: This attribute describes the firmware number of a datalogger. Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters long

Name: flags Relations: Station_Datalogger_LChannel Description: Channel flags. Datatype: char(27) NULL Value: Yes Range: Any combination of {T ¦ C ¦ H ¦ G ¦ W ¦ F ¦ S ¦ I ¦ E ¦ M ¦ B} T = channel is trigered C = channel is recorded continuously H = state of health data G = geophysical data W = weather or environmental data F = flag information (nominal, not ordinal) S = synthesized data I = channel is a calibration input E = channel is experimental or temporary M = maintenance tests are underway on channel; possible abnormal data B = data are a beam synthesis

Name: frequency Relations: Sensor_Component, Filter, Filter_Sequence, Filamp_PChannel Description: The frequency (fs) at which the value in sensitivity/gain is correct. Datatype: float NULL Value: Yes Range: frequency > 0.0 Units: Hertz

Name: gain Relations: Filamp_PChannel, Filter, Filter_Sequence, Filter_FIR Description: This attribute denotes the gain. Datatype: float NULL Value: Yes

Name: hp_id Relations: Response_HP Description: High-pass filter identifier. Each high-pass filter is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: i_error Relations: Response_PZ Description: The error value for the imaginary portion of a complex pole/zero. Datatype: float NULL Value: Yes

Name: i_value Relations: Response_PZ Description: The imaginary portion of a complex pole/zero. Datatype: float NULL Value: No

Name: in_sp_rate Relations: Filter Description: Sampling rate. This attribute is the input sample rate in samples/second. Datatype: float NULL Value: No Range: in_sp_rate > 0.0 Units: 1/seconds

Name: lat Relations: Station, Station_Sensor Description: Latitude. This attribute is the geographic latitude. Locations north of the equator have positive latitudes. Datatype: float NULL Value: Yes Range: -90.0  lat  90.0 Units: Degrees

Name: lchannel_nb Relations: Station_Datalogger_LChannel Description: This attribute denotes a logical channel of a datalogger. Datatype: int NULL Value: No Range: lchannel_nb  1

Name: lddate Relations: Response, Response_HP, Response_LP, Response_PZ, Response_PN, Sensor, Sensor_Component, Filamp, Filamp_PChannel, Datalogger, Datalogger_Board, Datalogger_Module, Filter_FIR, Filter, Filter_Sequence, Station, Station_Sensor, Station_Sensor_Component, Station_Filamp, Station_Filamp_PChannel, Station_Digitizer, Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: Load date. Date and time that the record was created or last modified, in Oracle date datatype. Datatype: date NULL Value: No Range: Any valid date between January 01, 4712 BC and January 01, 4712 AD Units: YYYY/MM/DD HH24:MI:SS

Name: location Relations: Station_Datalogger_LChannel Description: Describes the individual sites on an array station, operated by the same network operator. Datatype: char(2) NULL Value: Yes Range: Any free-format string up to 2 characters long

Name: lon Relations: Station, Station_Sensor Description: Longitude. This attribute is the geographic longitude. Longitudes are measured positive east of the Greenwich meridian. Datatype: float NULL Value: Yes Range: -180.0  lon  180.0 Units: Degrees

Name: lower_bound Relations: Response_PN Description: This attribute denotes the polynomial lower bound Datatype: float NULL Value: Yes

Name: lp_id Relations: Response_LP Description: Low-pass filter identifier. Each low-pass filter is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: max_error Relations: Response_PN Description: This attribute represents the maximum error of polynomial approximation. Datatype: float NULL Value: Yes

Name: module_nb Relations: Datalogger_Module Description: This attribute represents the module number of a datalogger’s board. Datatype: int NULL Value: No Range: module_nb  1

Name: name Relations: Response_PN, Sensor, Filamp, Filter_FIR, Filter_Sequence Description: This attribute denotes a descriptive name for a sensor, a filter-amplifier, a FIR filter, a chain of filters or a polynomial. Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters

Name: nb_aux_pchannel Relations: Station_Digitizer Description: This attribute denotes the number of auxiliary physical channels of a digitizer. Datatype: int NULL Value: No Range: nb_aux_pchannel  0

Name: nb_board Relations: Datalogger Description: This attribute denotes the number of boards of a datalogger. Datatype: int NULL Value: No Range: nb_board  0

Name: nb_coeff Relations: Response_PN Description: This attribute denotes the number of polynomial coefficients. Datatype: int NULL Value: No Range: nb_coeff  0

Name: nb_component Relations: Sensor, Station_Sensor Description: This attribute denotes the number of components of a sensor. Datatype: int NULL Value: No Range: nb_component  0

Name: nb_data Relations: Station Description: This attribute denotes the number of dataloggers of a station. Datatype: int NULL Value: No Range: nb_data  0

Name: nb_digi Relations: Station Description: This attribute denotes the number of digitizers of a station. Datatype: int NULL Value: No Range: nb_digi  0

Name: nb_filamp Relations: Station Description: This attribute denotes the number of filter-amplifiers of a station. Datatype: int NULL Value: No Range: nb_filamp  0

Name: nb_filter Relations: Filter_Sequence Description: This attribute denotes the number of filters of a sequence. Datatype: int NULL Value: No Range: nb_filter  0

Name: nb_lchannel Relations: Station_Datalogger_PChannel Description: This attribute denotes the number of logical channels of a datalogger. Datatype: int NULL Value: No Range: nb_lchannel  0

Name: nb_module Relations: Datalogger_Board Description: This attribute denotes the number of modules of a datalogger’s board. Datatype: int NULL Value: No Range: nb_module  0

Name: nb_pchannel Relations: Filamp, Station_Filamp, Station_Datalogger Description: This attribute denotes the number of physical channels of a datalogger. Datatype: int NULL Value: No Range: nb_pchannel  0

Name: nb_pole Relations: Response_HP, Response_LP Description: This attribute denotes the number of poles of a high/low-pass filter. Datatype: int NULL Value: No Range: nb_pole  0

Name: nb_pri_pchannel Relations: Station_Digitizer Description: This attribute denotes the number of primary physical channels of a digitizer. Datatype: int NULL Value: No Range: nb_pri_pchannel  0

Name: nb_sensor Relations: Station Description: This attribute denotes the number of sensors of a station. Datatype: int NULL Value: No Range: nb_sensor  0

Name: net Relations: Station, Station_Sensor, Station_Sensor_Component, Station_Filamp, Station_Filamp_PChannel, Station_Digitizer, Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: Unique network identifier. This character string is the name of a seismic network. Datatype: char(8) NULL Value: No Range: Any free-format string up to 8 characters long

Name: next_hard_nb Relations: Station_Sensor_Component, Station_Filamp_PChannel Description: This attribute denotes the next hardware number. Datatype: int NULL Value: No Range: next_hard_nb  1

Name: next_hard_pchannel Relations: Station_Sensor_Component, Station_Filamp_PChannel Description: This attribute denotes the next hardware physical channel. Datatype: int NULL Value: No Range: next_hard_pchannel  1

Name: next_hard_type Relations: Station_Sensor_Component, Station_Filamp_PChannel Description: This attribute denotes the next hardware type. Datatype: char(1) NULL Value: No Range: next_hard_type  { F ¦ D } F = filter-amplifier D = digitizer

Name: offdate Relations: Sensor, Filamp, Datalogger, Station, Station_Sensor, Station_Sensor_Component, Station_Filamp, Station_Filamp_PChannel, Station_Digitizer, Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: This attribute denotes the end date. Datatype: date NULL Value: Yes Range: Any valid date between January 01, 4712 BC and January 01, 4712 AD Units: YYYY/MM/DD HH24:MI:SS

Name: offset Relations: Filter Description: This field determines which sample is chosen for use. The value of this field has to be greater than or equal to zero, but less than the decimation factor. Datatype: int NULL Value: Yes Range: offset  0

Name: ondate Relations: Sensor, Filamp, Datalogger, Station, Station_Sensor, Station_Sensor_Component, Station_Filamp, Station_Filamp_PChannel, Station_Digitizer, Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: This attribute denotes the start date. Datatype: date NULL Value: Yes Range: Any valid date between January 01, 4712 BC and January 01, 4712 AD Units: YYYY/MM/DD HH24:MI:SS

Name: out_sp_rate Relations: Filter Description: Sampling rate. This attribute is the output sample rate in samples/second. Datatype: float NULL Value: No Range: out_sp_rate > 0.0 Units: 1/seconds

Name: pchannel_nb Relations: Filamp_PChannel, Station_Filamp_PChannel, Station_Digitizer_PChannel, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: This attribute denotes a physical channel. Datatype: int NULL Value: No Range: pchannel_nb  1

Name: pn_id Relations: Response_PN, Response_PN_Data Description: Polynomial identifier. Each polynomial is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: pn_nb Relations: Response_PN_Data Description: This attribute denotes the coefficient number of a polynomial. Datatype: int NULL Value: No Range: pn_nb  1

Name: pn_value Relations: Response_PN_Data Description: This attribute denotes the coefficient value of a polynomial. Datatype: float NULL Value: No

Name: poly_type Relations: Response_PN Description: This attribute represents the type of polynomial approximation. Datatype: char(1) NULL Value: No Range: poly_type  { C ¦ L ¦ M } C = Chebychev L = Legendre M = MacLaurin

Name: pz_id Relations: Response_PZ Description: Poles & zeros set identifier. Each poles & zeros set is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: pz_nb Relations: Response_PZ Description: This attribute denotes the pole/zero number of a response. Datatype: int NULL Value: No Range: pz_nb  1

Name: r_error Relations: Response_PZ Description: The error value for the real portion of a complex pole/zero. Datatype: float NULL Value: Yes

Name: r_type Relations: Response Description: A single character describing the type of response. Datatype: char(1) NULL Value: Yes Range: r_type  { A ¦ B ¦ C ¦ D} A = Analog (rad/sec) B = Analog (Hz) C = Composite D = Digital

Name: r_value Relations: Response_PZ Description: The real portion of a complex pole/zero. Datatype: float NULL Value: No

Name: remark Relations: Station_Datalogger_LChannel Description: An optional comment given to the instrument. It can be anything, including a serial number or a notation of special modifications. Datatype: char(30) NULL Value: Yes Range: Any free-format string up to 30 characters long

Name: resp_id Relations: Response Description: Response identifier. Each response is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: resp_nb Relations: Response Description: This attribute denotes the response number of a sequence. Datatype: int NULL Value: No Range: resp_nb  1

Name: resp_type Relations: Response Description: This attribute denotes the response type. Datatype: char(1) NULL Value: No Range: resp_type  { H ¦ L ¦ P ¦ Z ¦ F ¦ N } H = high-pass filter L = low-pass filter P = polynomial Z = poles and zeros F = FIR filter N = none

Name: rfrequency Relations: Station_Datalogger_LChannel Description: The frequency (fs) at which the value in rgain is correct. Datatype: float NULL Value: Yes Range: rfrequency > 0.0 Units: Hertz

Name: rgain Relations: Station_Datalogger_LChannel Description: This attribute denotes the gain at the response frequency. Datatype: float NULL Value: Yes

Name: samprate Relations: Station_Datalogger_LChannel Description: Sampling rate. This attribute is the sample rate in samples/second. Datatype: float NULL Value: No Range: samprate > 0.0 Units: 1/seconds

Name: seedchan Relations: Station_Datalogger_LChannel Description: SEED channel name. The first character denotes the band code, the second for the instrument code and the third represents the component code. Datatype: char(3) NULL Value: No Range: seedchan  {E ¦ S ¦ H ¦ B ¦ M ¦ L ¦ V ¦ U ¦ R}  {A ¦ B ¦ D ¦ F ¦ G ¦ H ¦ I ¦ K ¦ L ¦ M ¦ P ¦ R ¦ S ¦ V ¦ T ¦ W}  {Z ¦ N ¦ E ¦ A ¦ B ¦ C ¦ T ¦ R ¦ 1 ¦ 2 ¦ 3 ¦ U ¦ V ¦ W} Band code: Sample rate (Hz) Corner period (sec) E = extremely short period  80 < 10 S = short period  10 to < 80 < 10 H = high broad band  80  10 B = broad band  10 to < 80  10 M = mid period > 1 to < 10 L = long period ~ 1 V = very long period ~ 0.1 U = ultra long period ~ 0.01 R = extremely long period ~ 0.001 Instrument code: A = tiltmeter B = creepmeter D = pressure F = magnetometer G = gravimeter/accelerometer H = high gain seismometer I = humidity K = temperature L = low gain seismometer M = mass position seismometer P = geophone R = rainfall S = linear strainmeter V = volumetric strain T = tide meter W = wind Component code: N = north-south E = east-west T = transverse R = radial Z = vertical A, B, C = triaxial 1, 2, 3 = orthogonal components but non-traditional orientations U, V, W = optional components

Name: seed_io Relations: Station_Datalogger_PChannel Description: This attribute represents the SEED instrument and orientation. Datatype: char(2) NULL Value: No Range: seed_io  {A ¦ B ¦ D ¦ F ¦ G ¦ H ¦ I ¦ K ¦ L ¦ M ¦ P ¦ R ¦ S ¦ V ¦ T ¦ W}  {Z ¦ N ¦ E ¦ A ¦ B ¦ C ¦ T ¦ R ¦ 1 ¦ 2 ¦ 3 ¦ U ¦ V ¦ W} Instrument code: A = tiltmeter B = creepmeter D = pressure F = magnetometer G = gravimeter/accelerometer H = high gain seismometer I = humidity K = temperature L = low gain seismometer M = mass position seismometer P = geophone R = rainfall S = linear strainmeter V = volumetric strain T = tide meter W = wind Component code: N = north-south E = east-west T = transverse R = radial Z = vertical A, B, C = triaxial 1, 2, 3 = orthogonal components but non-traditional orientations U, V, W = optional components

Name: sensitivity Relations: Sensor, Datalogger_Module Description: This attribute denotes the sensitivity (Sd) for a sensor or a digitizer. Datatype: float NULL Value: No Units: Counts/Volt

Name: sensor_id Relations: Sensor, Sensor_Component, Station_Sensor Description: Sensor identifier. Each sensor is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: sensor_nb Relations: Station_Sensor, Station_Sensor_Component Description: This attribute represents the sensor number. Datatype: int NULL Value: No Range: sensor_nb  1

Name: seqfil_id Relations: Filter_Sequence, Filter_Sequence_Data, Station_Datalogger_LChannel Description: Sequence of filters identifier. Each sequence of filters is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: seqresp_id Relations: Response, Sensor_Component, Filamp_PChannel, Filter Description: Sequence of responses identifier. Each sequence of responses is assigned a unique integer which identifies it in a database. Datatype: int NULL Value: No

Name: serial_nb Relations: Sensor, Filamp, Datalogger, Datalogger_Board, Datalogger_Module, Station_Digitizer Description: This attribute denotes a serial number of a sensor, datalogger, filter-amplifier, board or module. Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters long

Name: software Relations Datalogger Description: Datalogger’s software (mshear, ushear, …). Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters long

Name: software_nb Relations Datalogger Description: Datalogger’s software version number. Datatype: char(80) NULL Value: Yes Range: Any free-format string up to 80 characters long

Name: sta Relations: Station, Station_Sensor, Station_Sensor_Component, Station_Filamp, Station_Filamp_PChannel, Station_Digitizer, Station_Digitizer_PChannel, Station_Datalogger, Station_Datalogger_PChannel, Station_Datalogger_LChannel Description: Station code. This is the common code-name of a seismic observatory. Generally only three or four characters are used. Datatype: char(6) NULL Value: No Range: Any free-format string up to 6 characters long

Name: staname Relations: Station Description: Station name/description. This is the full name of the station whose code-name is in sta. Datatype: char(50) NULL Value: Yes Range: Any free-format string up to 50 characters long

Name: symmetry Relations: Filter_FIR Description: This attribute denotes the eventual symmetry of a response. Datatype: char(1) NULL Value: No Range: symmetry  { E ¦ O ¦ N } E = even O = odd N = none

Name: type Relations: Response_PZ, Filter_FIR_Data Description: This attribute denotes whether a value is associated with a pole/zero or a numerator/denominator. Datatype: char(1) NULL Value: No Range: type  { P ¦ Z ¦ N ¦ D } P = pole Z = zero N = numerator D = denominator

Name: unit_calib Relations: Station_Datalogger_LChannel Description: A unit lookup key that refers to the Units Abbreviation Dictionary for the units of calibration input, usually volts or amps. Datatype: int NULL Value: Yes

Name: unit_in Relations: Response Description: A unit lookup key that refers to the Units Abbreviation Dictionary for the units of the incoming to this stage of the filter. This signal will usually be ground motion, volts, or counts, depending on where it is in the filter system. Datatype: int NULL Value: No

Name: unit_out Relations: Response Description: A unit lookup key that refers to the Units Abbreviation Dictionary for the stage’s output signal. Analog filters usually emit volts, and digital filters usually emit counts. Datatype: int NULL Value: No

Name: unit_signal Relations: Station_Datalogger_LChannel Description: A unit lookup key that refers to the Units Abbreviation Dictionary for the signal response of the instrument – usually the ground motion response, or something else for non-seismic devices. Datatype: int NULL Value: No

Name: upper_bound Relations: Response_PN Description: This attribute denotes the polynomial upper bound Datatype: float NULL Value: Yes

Name: word_16 Relations: Datalogger Description: 16-bit quantity byte swapping order, for the data headers only. Datatype: int NULL Value: No

Name: word_32 Relations: Datalogger Description: The swap order in which 32-bit quantities are specified in the data headers. Datatype: int NULL Value: No

A

azimuth	22

B

block_size	22

board_nb	22

board_type	23

C

channel	23

channel_comp	23

channel_type	24

channelsrc	24

clock_drift	25

coeff_nb	25

coefficient	25

comp_type	26

component_nb	26

component_type	26

corner_freq	27

correction	27

D

damping_value	27

data_format	28

data_id	28

data_nb	28

data_pchannel	29

data_type	29

Datalogger	6, 16

Datalogger_Board	6, 17

Datalogger_Module	6, 17

datumhor	29

datumver	30

delay	30

digi_channel	30

digi_nb	31

digi_polarity	31

digi_type	31

dip	32

E

edepth	32

elev	32

error	33

F

Filamp	5, 16

filamp_id	33

filamp_nb	33

Filamp_PChannel	5, 16

Filter	7, 18

Filter_FIR	7, 17

Filter_FIR_Data	7, 17

filter_id	33

filter_nb	34

Filter_Sequence	7, 18

Filter_Sequence_Data	8, 18

filter_type	34

fir_id	34

firmware_nb	35

flags	35

frequency	36

G

gain	36

H

hp_id	36

I

i_error	36

i_value	37

in_sp_rate	37

L

lat.	37

lchannel_nb	38

lddate	38

location	38

lon	39

lower_bound	39

lp_id	39

M

max_error	40

module_nb	40

N

name	40

nb_aux_pchannel	41

nb_board	41

nb_coeff	41

nb_component	42

nb_data	42

nb_digi	42

nb_filamp	43

nb_filter	43

nb_lchannel	43

nb_module	44

nb_pchannel	44

nb_pole	44

nb_pri_pchannel	45

nb_sensor	45

net	45

next_hard_nb	46

next_hard_pchannel	46

next_hard_type	46

O

offdate	47

offset	47

ondate	48

out_sp_rate	48

P

pchannel_nb	48

pn_id	49

pn_nb	49

pn_value	49

poly_type	50

pz_id	50

pz_nb	50

R

r_error	51

r_type	51

r_value	51

remark	52

resp_id	52

resp_nb	52

resp_type	53

Response	3, 14

Response_HP	3, 14

Response_LP	3, 14

Response_PN	4, 15

Response_PN_Data	4, 15

Response_PZ	4, 15

rfrequency	53

rgain	53

S

samprate	54

seed_io	56

seedchan	55

sensitivity	56

Sensor	4, 15

Sensor_Component	5, 16

sensor_id	57

sensor_nb	57

seqfil_id	57

seqresp_id	57

serial_nb	58

software	58

software_nb	58

sta.	59

staname	59

Station	8, 18

Station_Datalogger	10, 20

Station_Datalogger_LChannel	11, 21

Station_Datalogger_PChannel	11, 20

Station_Digitizer	10, 20

Station_Digitizer_PChannel	10, 20

Station_Filamp	9, 19

Station_Filamp_PChannel	9, 19

Station_Sensor	8, 19

Station_Sensor_Component	9, 19

symmetry	59

T

type	60

U

unit_calib	60

unit_in	60

unit_out	61

unit_signal	61

upper_bound	61

W

word_16	61

word_32	62

NCEDC Database Project Homepage: http://quake.geo.berkeley.edu/db/